A technical report of the additional test carried out to study the ProXR bias problem in Castilla y León (Spain)

a) Type of device tested and number of devices tested – Trimble ProXR; 3 devices were tested

b) Software (type and version) used – Trimble Pathfinder Office v.4.0

c) Firmware version – Asset Surveyor v.3.4 ; v4.03; v.5

d) Method of measurement – continuous, interval 2s

e) Device settings: max DOP 6, elevation mask 10º

f) Use of differential correction and type of correction – Two modes: standalone (no differential correction) and postprocessing with Pathfinder Office v.4.0

g) Number of sets and dates of measurement - 9 sets of measurements ; days 12, 13, 15, 18, 19 April 2011.

h) Number of repetitions per set - 4 repetitions , 2 in direction clokwise and 2 in direction anticlokwise.

i) List of operators –  3 , r=Rubén, j= Jose Carlos, m= Modesto

j) Number of parcels (No of parcels with trees on the border) - 1 parcel

Data:

k) Sketch of the test parcel on orthophotos and photos of the parcel - annex folder

l) The *shp files of the measurements - annex forder

m)  Measurements & statistics : Measurements_Statistics_ProXR_parcel33.xls

Motivation of the test

To try to clarify the issue of bias that showed up in the GPS validation test of Castilla y Leon (March 2011), we conducted in April a remeasurement test in 1 parcel, with thee ProXR receivers, one of them the same used in the original test (ProXR-2404), and two more (ProXR-3401, ProXR-4701). We made 9 sets, 4 repetitions per set, and also we compute in standalone mode and postprocessed mode as the original test.

Explanations of the validation process:

Reference area

The borders of the test field were marked with wooden pegs. The coordinates of each peg were measured with Leica GPS 1200 kit using the RTK method and VRS corrections from  GNSS network of Castilla y León (http://gnss.itacyl.es). The estimated precision is +- 2cm in planimetry. The coordinate system used was ED50-Spain.

Data collection with Trimble ProXR+Asset Surveyor

The data were collected in 9  independent sets with 4 repetitions on each set.

- recording data simultaneously by 3 operators, 3 sets per receiver per operator

- 2 sec between two points/continuous method

The data was collected in clockwise and counter clockwise direction to avoid systematic errors connected to left/right -handed operators and the direction of the walk. The measurements were made at different times of the day during 5 days.

Postprocessing of raw data

The raw data were also postprocessed (code) with Trimble Pathfinder Office v4 in order to compare conclusions with the original test for ProXR, both standalone an postprocessing.

# Statistical processing of the data

- Outliers detection:

- Visual Outliers: the measurements are uniform and not visual outliers in any receiver.

- Statistical tests: only 1 set excluded by Cochran’s test for ProXR-2404, standalone mode, set º 2.

- Bias significance test (Student’s test):

-ProXR-3401 & ProXR-4701, stand & postpro: no statistically significant bias at 1% level was noticed. The mean bias (diff (meas_area- ref_area)) is less than 10 m2.

-ProXR-2401, stand & postpro: no statistically significant bias at 1% level was noticed, but there is a high bias of around -230 m2 in the first 3 sets, and the following sets have a bias of only around -30 m2. It suggests us that perhaps an unnoticed change in the config parameters (we reset the receivers several times during the test) led to a fixing of the bias problem with this particular receiver, but we really don't know what it could be.

- Reproducibility limit:

The mean reproducibility limit, for the 3 receivers x 2 modes, is 0.67 m, slightly worse than the original test (0.58 m), but into the same buffer tolerance of 0.75 m.

# Conclusions

Our final conclusion is that it should always be carry out an area measurement in a reference parcel to assure the absence of significant bias, prior to first use, for those particular devices (Trimble ProXR).